## Fusion
last updated: | 2017-08-23
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
Fusion |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
Gases:
N2 | N2 gas @ 50psi, 1/8" tubing | Ultra high purity (UHP) >99.995% (5.0) from your local Gas supplier
N2 | N2 gas @ 100psi, 1/4" tubing | High purity (UHP) >99% from your local Gas supplier
He | He gas @ 40psi 1/8" copper tubing | Ultra high purity (UHP) >99.995% (5.0) from your local Gas supplier
Gas setup as per Thermo's preinstall guide:
Regulators and fittings, if the gas is delivered directly from cylinders
N2 100 psi | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 687.17 | 687.17 | Fisher
| SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 50 psi | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 494 | 494 | Restek | same regulator through Fisher: 06-711-718
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
He 40 psi | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 494 | 494 | Restek | same regulator through Fisher: 06-711-718
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
UWPR modified gas setup:
[8 empty] | UWPR Fusion |  |  |  | UWPR Lumos
NOTE we use NSI for the majority of time and only use the high flow source for calibrations, so our HP N2 consumption is relatively low. But with this setup we have the option to connect a LN2 dewar if needed.
UWPR uses 5.0 (99.999%)for all gases, NOTE tubing supplied by Thermo approx 6-8ft
UWPR uses 5.0 (99.999%) for both N2 supplies, we are stepping it down with a small inline regulator from 100psi to 50psi
Regulators and fittings, if the gas is delivered directly from cylinders
N2 both 50/100psi | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 687.17 | 687.17 | Fisher | you only need one regulator, the gas is split afterwards and the pressure stepped down to 50 psi with an inline regulator
| SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
He 40 psi | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 494 | 494 | Restek | same regulator through Fisher: 06-711-718
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
In line regulator | 22452 | Ultra-High Purity Chrome-Plated Brass Line Gas Regulator, 1/4″ female NPT ports, Outlet Pressure:0–100 psig | 1 | each | 315 | 315 | Restek | see Lumos picture
| 6763K82 | Miniature Air Regulator with Pressure Gauge Alum Body, 1/4 Pipe, Range 0-125 | 0 | each | 42.57 | 0 | McMaster Carr | cheaper alternative, see Fusion picture
| 1725K13 | High-Pressure 347 Stainless Steel Pipe Fitting, Plug with Hex Drive, 1/4 Anpt | 0 | each | 3.33 | 0 | McMaster Carr | you need one plug per regulator to replace the pressure relief valve
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.69 | 8.69 | Swagelok | In-line regulator exit to 1/8" tubing towards intrument
| SS-4-TA-1-4 | SS Swagelok Tube Fitting, Male Tube Adapter, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.31 | 6.31 | Swagelok | In-line regulator to tee
| SS-400-3 | SS Swagelok Tube Fitting, Union Tee, 1/4 in. Tube OD | 1 | each | 23.39 | 23.39 | Swagelok | Tee
| SS-43GXLS4 | SS 1-Piece 40G Series 3-Way Ball Valve, 0.90 Cv, 1/4 in. Swagelok Tube Fittings, L Flow Path | 1 | each | 127.41 | 127.41 | swagelok | 3-way L-valve
OR:
If the gas is delivered via a cylinder farm we use inline regulators to regulate each line to the instrument, as opposed to the cylinder regulators
N2 both 50/100psi | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | use inline regulators if the gas is delivered via a cylinder farm
| SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | fitting regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
He 40 psi | 22452 | Ultra-High Purity Chrome-Plated Brass Line Gas Regulator, 1/4″ female NPT ports, Outlet Pressure:0–100 psig | 1 | each | 235.87 | 235.87 | Restek | Fisher part # 06-719-402
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | fitting regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
In line regulator | 22452 | Ultra-High Purity Chrome-Plated Brass Line Gas Regulator, 1/4″ female NPT ports, Outlet Pressure:0–100 psig | 1 | each | 315 | 315 | Restek | see Lumos picture
| 6763K82 | Miniature Air Regulator with Pressure Gauge Alum Body, 1/4 Pipe, Range 0-125 | 0 | each | 42.57 | 0 | McMaster Carr | cheaper alternative, see Fusion picture
| 1725K13 | High-Pressure 347 Stainless Steel Pipe Fitting, Plug with Hex Drive, 1/4 Anpt | 0 | each | 3.33 | 0 | McMaster Carr | you need one plug per regulator to replace the pressure relief valve
| SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.69 | 8.69 | Swagelok | In-line regulator exit to 1/8" tubing towards intrument
| SS-4-TA-1-4 | SS Swagelok Tube Fitting, Male Tube Adapter, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.31 | 6.31 | Swagelok | In-line regulator to tee
| SS-400-3 | SS Swagelok Tube Fitting, Union Tee, 1/4 in. Tube OD | 1 | each | 23.39 | 23.39 | Swagelok | Tee
| SS-43GXLS4 | SS 1-Piece 40G Series 3-Way Ball Valve, 0.90 Cv, 1/4 in. Swagelok Tube Fittings, L Flow Path | 1 | each | 127.41 | 127.41 | swagelok | 3-way L-valve
LN2 dewar | 10-575-137 | Fisherbrand™ Multistage Cylinder Regulators |  | each | 775 | 0 | Fisher | optional regulator if you want to connect an LN2 Dewar
Instrumentation quick connects to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT |  | each | 17.02 | 0 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Power | Please consult the pre-install guide and the installing service engineer
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
230 VAC | fused with 20 A (15 or 16 A based on the preinstall guide, but local code may require 20 A) |  | 1
Note the power cord comes with a 16 A version of a 3 pole CEE male connector
Our electrician installed a NEMA 630 outlet and provided the corresponding NEMA plug
208 VAC | for the vacuum pump |  | 1 | for the Fusion
| Note the power cord comes with a 16 A version of a 3 pole CEE male connector |  | 2 | for the Lumos
Our electrician installed a NEMA 630 outlet and provided the corresponding NEMA plug
Alternatively Thermo sells a Power Var that will use 208 VAC and power both the instrument and the pump.. Check with your sales rep or local engineer.
Table
here is what we are using:
| WB250112 | MOBILE COMPUTER WORKCENTER WITH SLIDE-OUT KEYBOARD/MOUSE DRAWER | 1 | each | 174.95 | 174.95 | Global Industrial |  |  |  | WB250112
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the lower platform
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | most drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup
MS N2 100 psi
MS N2 50 psi
3-way L-valve
MS N2 50 psi
3-way L-valve
from the cyliner
MS N2 100 psi
MS N2 100 psi
MS N2 50 psi
3-way L-valve
open, optional
connect to LN2 dewar
Open
or connect to LN2 dewar
To the back of the MS
N2 100 psi in
To the back of the MS
N2 50 psi in
¼” tubing
¼” tubing
1/8” tubing
¼” tubing
¼” tubing
From gas cylinder
N2 at 100 psi
3-Way Ball Valve
1/4 in. Tube OD
SS-43GXLS4
Tee,
1/4 in. Tube OD
SS-400-3
Male Tube Adapter
1/4 in. Male NPT x
1/4 in. Tube OD
SS-4-TA-1-4
Male Connector,
1/8 in. Tube OD x
1/4 in. Male NPT
SS-200-1-4
Line Gas Regulator
Ultra-High Purity
22452
Cheaper option
Air Regulator
6763K82
Requires Plug
1/4 NPT
1725K13
10-575-137

## QExactive
last updated: | 2014-04-22
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
Q Exactive
Q Exactive plus |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
Gas:
N2 | N2 gas @ 116 psi (+/- 4 psi), 1/4" tubing | Ultra high purity (UHP) >99.995% (5.0) from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 687.17 | 687.17 | Fisher
Fittings for regulators
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | If the gas is delivered via a cylinder farm |  |  | no fisher part # available | for N2 gas @ 100psi
Fittings for regulators
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Power | Please consult the pre-install guide and the installing service engineer
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
230 VAC | fused with 20 A (15 or 16 A based on the preinstall guide, but local code may require 20 A) |  | 1
Note the power cord comes with a 16 A version of a 3 pole CEE male connector
Our electrician installed a NEMA 630 outlet and provided the corresponding NEMA plug
208 VAC | for the vacuum pump |  | 1
Note the power cord comes with a 16 A version of a 3 pole CEE male connector
Our electrician installed a NEMA 630 outlet and provided the corresponding NEMA plug
Alternatively Thermo sells a Power Var that will use 208 VAC and power both the instrument and the pump.. Check with your sales rep or local engineer.
Table
here is what we are using:
| 5' Series 1000 Mobile Table | 36" D x 60" L x 30-37" Adjustable Work height, Grey Glace Plastic Laminate Tops, #0001 Tan powder-Coat Paint finish | 1 |  |  |  |  | Note we also use 6ft tables, all depends on space constraints you're working with
| 8933 | Top Gray Glace 36DX60L | 1 | each | 167.2 | 167.2 | WorkplaceSystems
| 1013-0001 | Frame 1000 36DX60L | 1 | each | 64.9 | 64.9 | WorkplaceSystems
| 1164SE-0001 | Sway brace 1000 60L | 1 | each | 14.4 | 14.4 | WorkplaceSystems
| 1015-0001 | End caps plain 2x2 | 4 | each | 1.5 | 6 | WorkplaceSystems
| 1234M-0001 | Leg 1000MBL 36DX30/37.5H | 2 | each | 96.4 | 192.8 | WorkplaceSystems
| 2012-01 | Caster 5x1-1/4 Total lock BB wheel, Poly, 250# cap | 4 | each | 31.5 | 126 | WorkplaceSystems
| 1279-0001 | Mobile leg sway bar - 60L (51.852L W/2 BRKTS) stock | 1 | each | 37.9 | 37.9 | WorkplaceSystems
| N3115-0001 | Drawer case 2-3109, 1-3110 stock | 1 | each | 206.7 | 206.7 | WorkplaceSystems | Note drawers are optional, but nice to have
| 3123HD | Drawer brackets 36" heavy duty | 1 | each | 15.9 | 15.9 | WorkplaceSystems
831.8
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the lower platform on the legs
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | most drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup

## LTQ-OT install
last updated: | 2009-10-26
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
LTQ-OT
LTQ-OT XL |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
Gas:
N2 | N2 gas @ 100psi, 1/4" tubing | Ultra high purity >99.995% from your local Gas supplier |  |  |  |  | Polar cryogenics | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
He | He gas @ 40psi 1/8" copper tubing | Ultra high purity >99.995% from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 802.37 | 802.37 | Fisher
He | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 463 | 463 | Restek
He | 06-711-718 | Restek* Ultra-High-Purity Brass Body Regulators (Restek# 21667) |  |  | 475 | 0 | Fisher | same regulator through Fisher
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | If the gas is delivered via a cylinder farm |  |  | no fisher part # available | for N2 gas @ 100psi
He | 81891 | Single-Stage Regulator, delivery range 0-120 psi, gauge 0-200 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | purchase path |  |  | no fisher part # available | for He gas @ 40psi
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Distilled Water:
| 23-749999 | Water; deionized; distilled; Type II; HARLECO; ; H2O, Cubitainer; 20L | 1 | each | 60.9 | 60.9 | Fisher | Thermo recommends to use distilled water to be used in the water chiller, to minimize corrosion cause by disimilar metals.
Power | Please consult the pre-install guide and the installing service engineer
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
3 x 220 VAC | three phase 5wire system in Y configuration |  | 1
Note if you will be using the Power Var follow specific instructions related to that.
We had a transformer and outlet installed by an electrician, they supplied us with an US standard twist lock plug as well.
Table
here is what we are using:
| WB250112 | MOBILE COMPUTER WORKCENTER WITH SLIDE-OUT KEYBOARD/MOUSE DRAWER | 1 | each | 174.95 | 174.95 | Global Industrial |  |  |  | WB250112
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the lower platform
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
Noise reduction boxes for the two rough pumps:
| OrbiBox | Noise Reduction Box including all necessary parts to dampen both pumps in the ThermoFisher OrbiTrap | 1 | each | 2300 | 2300 | Sonation GmbH | http://www.sonation.de/pages/en/home.php
| SSH56TF | Noise Reduction Box for waterchillers like the NESLAB ThermoFlex 900/1400 | 1 | each | 2300 | 2300 | Sonation GmbH
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | Drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup

## LTQ-FT install
last updated: | 2009-10-26
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
LTQ-FT
LTQ-FT Ultra |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
Gas:
N2 | N2 gas @ 100psi, 1/4" tubing | Ultra high purity >99.995% from your local Gas supplier |  |  |  |  | Polar cryogenics | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
He | He gas @ 40psi 1/8" copper tubing | Ultra high purity >99.995% from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 802.37 | 802.37 | Fisher
He | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 463 | 463 | Restek
He | 06-711-718 | Restek* Ultra-High-Purity Brass Body Regulators (Restek# 21667) |  |  | 475 | 0 | Fisher | same regulator through Fisher
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | If the gas is delivered via a cylinder farm |  |  |  | no fisher part # available | for N2 gas @ 100psi
He | 81891 | Single-Stage Regulator, delivery range 0-120 psi, gauge 0-200 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | purchase path |  |  |  | no fisher part # available | for He gas @ 40psi
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Cryogens:
LN2 |  | LN2 500 L for install (2x240 L dewars) and ~240 L every 14 days |  |  | check with your local vendor |  | Polar cryogenics | we have an Oxford Magnet, numbers may vary for Jastek magnet
LHe |  | LHe 500 L for install (2x250 dewars) and ~100 L approx 6 times a year |  |  | check with your local vendor
Goddard fitting for LHe dewar, down to 3/8 in |  | Fittings for He dewar (Goddard fittings will be required for the Helium Dewar. Not all Dewars will arrive with these fittings so you should have them in the laboratory. Fittings that step down to a 3/8” diameter are recommended (typically this will be four fittings for a 250 L Dewar) 250 L dewars have 3/4" neck opening and 100 L dewars have 1/2" opening) Note any cryogenic supplier should have those.
Oxford magnet transfer tube requires a 3/8” Goddard fitting. | 3A-2000-5B | Brass solid end plug for 3/8" Goddard | 2 | each | 11 | 22 | Macomber Cryogenics | (508) 981-5632 Harry Schaller (Sales) | harry@macombercryogenics.com
| B-38-M | Brass Goddard vacuum coupling tube OD 3/8" | 2 | each | 34 | 68 | Macomber Cryogenics |  | http://www.goddardvalve.com/quikconn.htm
| B-50-M | Brass Goddard vacuum coupling tube OD 1/2" | 2 | each | 34 | 68 | Macomber Cryogenics
| B-62-M | Brass Goddard vacuum coupling tube OD 5/8" | 2 | each | 64 | 128 | Macomber Cryogenics | only on 250L Dewars
| B-75-M | Brass Goddard vacuum coupling tube OD 3/4" | 2 | each | 64 | 128 | Macomber Cryogenics | only on 250L Dewars
LN2 transfer tube | CFUL-6-.5 | 1/2" female flare one end. 3/8"o.d. tube, 90 degrees x 24". | 2 | each | 185 | 370 | Cryofab | Vincent Grillo, 908 686 3636 X15 |  | vgrillo@cryofab.com | see figure below
to connect LN2 transfer line that comes with Oxford magnet to LN2 dewar: | [7 empty] | Note: we used both the LN2 transfer tube from Cryofab as well as the one that came with the magnet, so we could fill both the LN2 and the Lhe tank with LN2 during the install.
| SS-810-A-8ANF | SS Swagelok Tube Fitting, Female Adapter, 1/2 in. Tube OD x 1/2 in. AN Tube Flare | 1 | each | 14.2 | 14.2 | Swagelok
| SS-810-6 | SS Swagelok Tube Fitting, Union, 1/2 in. Tube OD | 3 | each | 19.8 | 59.4 | Swagelok
| T-810-SET | PTFE Ferrule Set (1 Front Ferrule/1 Back Ferrule) for 1/2 in. Swagelok Tube Fitting | 10 | each | 3.05 | 30.5 | Swagelok
Helium “pusher”
to connect He pusher to He source | EF5383N | Daigger® PFA Tubing 1/4" OD | 1 | pack (50ft) | 76 | 76 | Daigger | http://www.daigger.com/
to connect 1/4" tubing to He pusher | SS-QC4-D-400 | SS Instrumentation Quick-Connect Stem w/ Valve, 0.2 Cv, 1/4 in. Swagelok Tube Fitting | 1 | each | 23.701 | 23.701 | Swagelok | Reed Perry, 425.531.0366 cell; 425.825.1115 office; 425.825.1705 fax | Reed.Perry@Seattle.swagelok.com
| SS-QC4-B-4PM | SS Instrumentation Quick-Connect Body, 0.2 Cv, 1/4 in. Male NPT | 1 | each | 26.2 | 26.2 | Swagelok | Seattle Fluid System Technologies | http://www.swagelok.com/
| KCP1DLC2A2P100J0-AS | SS Compact Pressure-Reducing Regulator, 25 psig | 1 | each | 285 | 285 | Swagelok
| SS-4-HRN-2 | Stainless Steel Pipe Fitting, Hex Reducing Nipple, 1/4 in. Male NPT x 1/8 in. Male NPT | 1 | each | 5.1 | 5.1 | Swagelok |  | complete instrumentation quick connect catalog
| SS-4-T | Stainless Steel Pipe Fitting, Tee, 1/4 in. Female NPT | 1 | each | 19.8 | 19.8 | Swagelok
| SS-4P4T2 | SS Quarter-Turn Instrument Plug Valve, 1/4 in. MNPT, 1.0 Cv | 1 | each | 40.3 | 40.3 | Swagelok
| SS-4-HLN-4.00 | Stainless Steel Pipe Fitting, Hex Long Nipple, 1/4 in. Male NPT, 4 in. Length | 1 | each | 12.1 | 12.1 | Swagelok
| SS-4-HCG | Stainless Steel Pipe Fitting, Hex Coupling, 1/4 in. Female NPT | 1 | each | 8.2 | 8.2 | Swagelok
| SS-QC4-B-4PM | SS Instrumentation Quick-Connect Body, 0.2 Cv, 1/4 in. Male NPT | 1 | each | 26.2 | 26.2 | Swagelok
| SS-QC4-S-4PM | stem, 1/4 male NPT, w/o shutoff | 1 | each | 14.1 | 14.1 | Swagelok
to connect pusher to dewar (which has 1/2" female NPT) | SS-8-RB-4 | Stainless Steel Pipe Fitting, Reducing Bushing, 1/2 in. Male NPT x 1/4 in. Female NPT | 1 | each | 9.2 | 9.2 | Swagelok
Note: the pressure regulator and quickconnects are optional
Note: for a more economical version you can also use brass instead of SS
Regulator (range up to 60 psi, will be used to push LHe out of the dewar typically about 10 psi, 1/4" swagelok)(Regulator for Helium tank to pressurize the Helium Dewar. The regulator should have a 1/4” fitting on the output and be low pressure (0-15 psi)) | 06-710-882 | Restek* Ultra-High-Purity Stainless-Steel Body Regulators, He 0-100psi outlet to 1/4"FTP outlet, dual stage | 1 | each | 666.9 | 666.9 | Fisher | Note next time I would get one with smaller outlet gauge range, but this one worked
Power | Please consult the pre-install guide and the installing service engineer
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
3 x 220 VAC | three phase 5wire system in Y configuration |  | 1
Note if you will be using the Power Var follow specific instructions related to that.
We had a transformer and outlet installed by an electrician, they supplied us with an US standart twist lock plug as well.
Table | here is what we are using:
| WB250112 | MOBILE COMPUTER WORKCENTER WITH SLIDE-OUT KEYBOARD/MOUSE DRAWER | 1 | each | 174.95 | 174.95 | Global Industrial |  |  |  |  | WB250112
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the lower platform
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
| 59580-386 | VWR* Black Rubber Stoppers, Solid | 1 | pack | 28.09 | 28.09 | VWR | stoppers #10
| S47610 | Absorbent Lab Towels | 1 | pack | 16.65 | 16.65 | Fisher | Towels.
| WB253939 | 6' Fiberglass Platform Step Ladder - 300 Lb Capacity | 1 | each | 189.95 | 189.95 | Global Industrial | Six foot step ladder (non magnetic)
| EF978BX | cryo gloves medium | 1 | pair | 146 | 146 | Daigger | Thermal gloves. Note you'll need gloves for two people for LHe fills.
| EF978CX | cryo gloves large | 1 | pair | 146 | 146 | Daigger | Thermal gloves. Note you'll need gloves for two people for LHe fills.
| 5490T56 | Size Medium: Premium All Leather Cold-Protection Glove Grain Deerskin, Slip On, 10" Length, Gold | 1 | each | 29.65 | 29.65 | Mcmaster Carr | Leather gloves for better dexterity. Note you'll need gloves for two people for LHe fills.
| 5490T56 | Size Large: Premium All Leather Cold-Protection Glove Grain Deerskin, Slip On, 10" Length, Gold | 1 | each | 29.65 | 29.65 | Mcmaster Carr | Leather gloves for better dexterity. Note you'll need gloves for two people for LHe fills.
Tools: | Note it helps to have a reasonable set of tools at hand
| 941206000 | Craftsman 96 pc. Field Technicians Tool Set | 1 | set | 179.99 | 179.99 | Sears | Allen wrenches – Inch and metric
| NC9014921 | Adjustable Wrenches, 10 in. , Opening Size: 1 1/8 in. , Finish/Coating: Chrome. | 2 | each | 25.19 | 50.38 | Fisher | Adjustable wrenches (2) (10”).
| 5368A18 | V-Jaw Tongue-and-Groove Pliers Cushion-Grip, 16" L, 4-1/2" Max Opening, 8 Positions | 1 | each | 29 | 29 | Mcmaster Carr | Pliers
| 26408-094 | HEAT GUN 500-650F 5AMP 120V | 1 | each | 127.92 | 127.92 | VWR | Heat gun
| NC9763786 | Cord, Extension, Rugged Outdoor, 12-Gauge Wire, 15A, 1875w, Three-Wire Grounded, | 1 | each | 27.47 | 27.47 | Fisher | Extension cord (14 gauge, grounded, 25’ or as is appropriate in your laboratory).
Backup O-rings | 9557K475 | EPDM O-ring ODxIDxWidth 1/2x3/8x2/32 pk of 100 | 1 | each | 11.15 | 11.15 | Mcmaster Carr | O-ring for Goddard fitting
| 9557K466 | EPDM O-ring ODxIDxWidth 5/8x1/2x2/32 pk of 100 | 1 | each | 7.53 | 7.53 | Mcmaster Carr | O-ring for Goddard fitting
| 9557K464 | EPDM O-ring ODxIDxWidth 13/16x5/8x3/32 pk of 100 | 1 | each | 6.23 | 6.23 | Mcmaster Carr | O-ring for Goddard fitting |  | only on 250L Dewars
| 9557K477 | EPDM O-ring ODxIDxWidth 15/16x3/4x3/32 pk of 50 | 1 | each | 6.26 | 6.26 | Mcmaster Carr | O-ring for Goddard fitting |  | only on 250L Dewars
| 9557K471 | EPDM O-ring ODxIDxWidth 9/16x3/8x3/32 pk of 100 | 1 | each | 9.48 | 9.48 | Mcmaster Carr | O-ring for LN2 transfer line
Replacement transfer tube entry port on Oxford magnet (for 3/8" transfer tubes, LN2 and LHe) | [7 empty] | Oxford Instruments America, Inc. 130A Baker Ave. Ext. Concord, MA 01742
| A2-101 | Transfer tube entry port, 0.375 inches | 2 | each | 79.58 | 159.16 | Oxford | Note: it is nice to have a second set, so you can leave one attached to the transfer line and the other with the magnet. You'll know what I mean when you do your first fill.
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | Drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup
LN2 transfer tube | CFUL-6-.5

## TSQ install
last updated: | 2009-10-26
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
TSQ Quantum Ultra
TSQ Quantum Access |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
TSQ Vantage
Gas:
N2 | N2 gas @ 100psi, 1/4" tubing | high purity >99% from your local Gas supplier |  |  |  |  | Polar cryogenics | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Ar | Ar gas @ 20psi 1/8" copper tubing | Ultra high purity >99.995% from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 802.37 | 802.37 | Fisher
Ar | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 463 | 463 | Restek | NOTE there might be others with a smaller outlet pressure range, e.g. 0-50 psi or so….you'd have to check.
Ar | 06-711-718 | Restek* Ultra-High-Purity Brass Body Regulators (Restek# 21667) |  |  | 475 | 0 | Fisher | same regulator through Fisher
Fittings for regulators
Ar | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech
Ar | 81890 | single stage regulator 0-40 psi | 1 | each | 247.66 | 247.66 | Alltech
| 06-711-717 | Restek* Ultra-High-Purity Brass Line Regulator 0-50 psi, ports 1/4" FPT (resteck # 21666) | 1 | each | 260.1 | 260.1 | Fisher | Alternate
Fittings for regulators
Ar | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
TSQ Vantage | you'll need a split for the N2 line with an extra regulator:
| 06-710-846 | Restek MINICYL Regulator (20611), 0-60psig, 1/4 in. fittings | 1 | each | 55.18 | 55.18 | Fisher
| SS-400-3 | SS Swagelok Tube Fitting, Union Tee, 1/4 in. Tube OD | 1 | each | 22.35 | 22.35 | Swagelok
Power
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
220 VAC | 220 VAC | Note 208VAC is not enough, you need to boost it to min 215 VAC | 1
| Outlet: | NEMA 6-15 or NEMA 6-20 both will work
the power cord to the instrument is about 6ft long
| Breaker | 20 Amp (min 16 Amp)
Table
here is what we are using:
| 5' Series 1000 Mobile Table | 36" D x 60" L x 30-37" Adjustable Work height, Grey Glace Plastic Laminate Tops, #0001 Tan powder-Coat Paint finish | 1 |  |  |  |  | Note we also use 6ft tables, all depends on space constraints you're working with
| 8933 | Top Gray Glace 36DX60L | 1 | each | 167.2 | 167.2 | WorkplaceSystems
| 1013-0001 | Frame 1000 36DX60L | 1 | each | 64.9 | 64.9 | WorkplaceSystems
| 1164SE-0001 | Sway brace 1000 60L | 1 | each | 14.4 | 14.4 | WorkplaceSystems
| 1015-0001 | End caps plain 2x2 | 4 | each | 1.5 | 6 | WorkplaceSystems
| 1234M-0001 | Leg 1000MBL 36DX30/37.5H | 2 | each | 96.4 | 192.8 | WorkplaceSystems
| 2012-01 | Caster 5x1-1/4 Total lock BB wheel, Poly, 250# cap | 4 | each | 31.5 | 126 | WorkplaceSystems
| 1279-0001 | Mobile leg sway bar - 60L (51.852L W/2 BRKTS) stock | 1 | each | 37.9 | 37.9 | WorkplaceSystems
| N3115-0001 | Drawer case 2-3109, 1-3110 stock | 1 | each | 206.7 | 206.7 | WorkplaceSystems | Note drawers are optional, but nice to have
| 3123HD | Drawer brackets 36" heavy duty | 1 | each | 15.9 | 15.9 | WorkplaceSystems
831.8
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| WB190320 | Power Strip 6 Outlets And Surge Protector, 6 ft cord | 0 | each | 44.95 | 0 | Global Industrial
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the platform of the leg
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
Noise reduction boxes for the two rough pumps: |  |  | NOTE the TSQ Quantum Access only has one rough pump
| SSH11TF | Noise Reduction Box SSH11TF 350x300x750 (WxHxD) | 2 | each | 1270 | 2540 | Sonation GmbH | http://www.sonation.de/pages/en/home.php
| OP1141 | Oil pan for SSH11TF | 2 | each | 240 | 480 | Sonation GmbH
| FK-KF25 | Flange Extender Kit consisting of elongation 100mm, gasket and clamp | 4 | each | 48 | 192 | Sonation GmbH
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | Drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup

## LTQ install
last updated: | 2009-10-26
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
LTQ
LTQ XL |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
LTQ Velos
Gas:
N2 | N2 gas @ 100psi, 1/4" tubing | high purity >99% from your local Gas supplier |  |  |  |  | Polar cryogenics | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
He | He gas @ 40psi 1/8" copper tubing | Ultra high purity >99.995% from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 | each | 802.37 | 802.37 | Fisher
He | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 463 | 463 | Restek
He | 06-711-718 | Restek* Ultra-High-Purity Brass Body Regulators (Restek# 21667) |  |  | 475 | 0 | Fisher | same regulator through Fisher
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | If the gas is delivered via a cylinder farm |  |  | no fisher part # available | for N2 gas @ 100psi
He | 81891 | Single-Stage Regulator, delivery range 0-120 psi, gauge 0-200 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | purchase path |  |  | no fisher part # available | for He gas @ 40psi
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Power
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
220 VAC | 220 VAC | Note 208VAC is not enough, you need to boost it to min 215 VAC | 1
| Outlet: | NEMA 6-15 or NEMA 6-20 both will work
the power cord to the instrument is about 6ft long
| Breaker | 20 Amp (min 16 Amp)
Table
here is what we are using:
| 5' Series 1000 Mobile Table | 36" D x 60" L x 30-37" Adjustable Work height, Grey Glace Plastic Laminate Tops, #0001 Tan powder-Coat Paint finish | 1 |  |  |  |  | Note we also use 6ft tables, all depends on space constraints you're working with
| 8933 | Top Gray Glace 36DX60L | 1 | each | 167.2 | 167.2 | WorkplaceSystems
| 1013-0001 | Frame 1000 36DX60L | 1 | each | 64.9 | 64.9 | WorkplaceSystems
| 1164SE-0001 | Sway brace 1000 60L | 1 | each | 14.4 | 14.4 | WorkplaceSystems
| 1015-0001 | End caps plain 2x2 | 4 | each | 1.5 | 6 | WorkplaceSystems
| 1234M-0001 | Leg 1000MBL 36DX30/37.5H | 2 | each | 96.4 | 192.8 | WorkplaceSystems
| 2012-01 | Caster 5x1-1/4 Total lock BB wheel, Poly, 250# cap | 4 | each | 31.5 | 126 | WorkplaceSystems
| 1279-0001 | Mobile leg sway bar - 60L (51.852L W/2 BRKTS) stock | 1 | each | 37.9 | 37.9 | WorkplaceSystems
| N3115-0001 | Drawer case 2-3109, 1-3110 stock | 1 | each | 206.7 | 206.7 | WorkplaceSystems | Note drawers are optional, but nice to have
| 3123HD | Drawer brackets 36" heavy duty | 1 | each | 15.9 | 15.9 | WorkplaceSystems
831.8
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| WB190320 | Power Strip 6 Outlets And Surge Protector, 6 ft cord | 0 | each | 44.95 | 0 | Global Industrial
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the platform of the leg
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
Noise reduction boxes for the two rough pumps:
| SSH11TF | Noise Reduction Box SSH11TF 350x300x750 (WxHxD) | 2 | each | 1270 | 2540 | Sonation GmbH | http://www.sonation.de/pages/en/home.php
| OP1141 | Oil pan for SSH11TF | 2 | each | 240 | 480 | Sonation GmbH
| FK-KF25 | Flange Extender Kit consisting of elongation 100mm, gasket and clamp | 4 | each | 48 | 192 | Sonation GmbH
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | Drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup

## LTQ-ETD install
last updated: | 2009-10-26
These are suggestions of parts/equipment you likely will need when installing a new mass spectrometer. This list is based on things we are using at the UWPR.
Please consult the Thermo Pre-install guide specific to your instrument for detailed information! You can request one from your sales rep.
| part # | Description | Qty | unit | unit price | total | vendor | comment
LTQ-ETD |  | Number of instruments to be installed: | 1 |  |  |  |  | Change this number to adjust the number of parts you need below!
Gas:
N2 | N2 gas @ 100psi, 1/4" tubing | high purity >99% from your local Gas supplier |  |  |  |  | Polar cryogenics | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
He | He gas @ 40psi 1/8" copper tubing | Ultra high purity >99.995% from your local Gas supplier | [5 empty] | UWPR uses 5.0 (99.999%), NOTE tubing supplied by Thermo approx 6-8ft
Regulators if the gas is delivered directly from a cylinder
N2 | 10-575-137 | Fisherbrand* Multistage Cylinder Regulators, CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 400psig, Inlet pressure max 3000 psi, Outlet pressure 0-250 psi | 1 |  | 802.37 | 802.37 | Fisher
He | 21667 | Dual-Stage Ultra-High-Purity Chrome-Plated Brass Regulators CGA 580, Inlet gauge: 0 to 4,000psig, Outlet gauge: 0 to 200psig, Inlet pressure max 3000 psi, Outlet pressure 0-100 psi | 1 | each | 463 | 463 | Restek
He | 06-711-718 | Restek* Ultra-High-Purity Brass Body Regulators (Restek# 21667) |  |  | 475 | 0 | Fisher | same regulator through Fisher
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
OR:
If the gas is delivered via a cylinder farm
In line regulators
N2 | 81892 | Single-Stage Regulator, delivery range 0-200 psi, gauge 0-400 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | If the gas is delivered via a cylinder farm |  |  | no fisher part # available | for N2 gas @ 100psi
He | 81891 | Single-Stage Regulator, delivery range 0-120 psi, gauge 0-200 psi, 1/4" FPT ports | 1 | each | 235.87 | 235.87 | Alltech | purchase path |  |  | no fisher part # available | for He gas @ 40psi
Fittings for regulators
He | SS-200-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/8 in. Tube OD x 1/4 in. Male NPT | 1 | each | 8.3 | 8.3 | Swagelok | regulator out to He line (1/8'): you need a fitting from the regulator to 1/8" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
N2 | SS-400-1-4 | SS Swagelok Tube Fitting, Male Connector, 1/4 in. Tube OD x 1/4 in. Male NPT | 1 | each | 6.46 | 6.46 | Swagelok | regulator out to N2 line (1/4'): you need a fitting from the regulator to 1/4" tube fitting, most regulators have a 1/4" FPT, but double check the fitting on the regulator you get
Fitting to connect to your house line
| SS-QC4-S-4PM | SS Instrumentation Quick-Connect Stem w/o Valve, 0.3 Cv, 1/4 in. Male NPT | 2 | each | 17.02 | 34.04 | Swagelok | quick connect to regulator inlet, this is what we use at UWPR, but you need to find the correct fitting for your house line
Power
110 VAC | outlets to plug in your PC, Monitor, Switchbox and other accessories, like HPLC etc… |  | 4
220 VAC | 220 VAC | Note 208VAC is not enough, you need to boost it to min 215 VAC | 2
| Outlet: | NEMA 6-15 or NEMA 6-20 both will work
the power cord to the instrument is about 6ft long
| Breaker | 20 Amp (min 16 Amp)
Table
here is what we are using:
| 5' Series 1000 Mobile Table | 36" D x 60" L x 30-37" Adjustable Work height, Grey Glace Plastic Laminate Tops, #0001 Tan powder-Coat Paint finish | 1 |  |  |  |  | Note we also use 6ft tables, all depends on space constraints you're working with
| 8933 | Top Gray Glace 36DX60L | 1 | each | 167.2 | 167.2 | WorkplaceSystems
| 1013-0001 | Frame 1000 36DX60L | 1 | each | 64.9 | 64.9 | WorkplaceSystems
| 1164SE-0001 | Sway brace 1000 60L | 1 | each | 14.4 | 14.4 | WorkplaceSystems
| 1015-0001 | End caps plain 2x2 | 4 | each | 1.5 | 6 | WorkplaceSystems
| 1234M-0001 | Leg 1000MBL 36DX30/37.5H | 2 | each | 96.4 | 192.8 | WorkplaceSystems
| 2012-01 | Caster 5x1-1/4 Total lock BB wheel, Poly, 250# cap | 4 | each | 31.5 | 126 | WorkplaceSystems
| 1279-0001 | Mobile leg sway bar - 60L (51.852L W/2 BRKTS) stock | 1 | each | 37.9 | 37.9 | WorkplaceSystems
| N3115-0001 | Drawer case 2-3109, 1-3110 stock | 1 | each | 206.7 | 206.7 | WorkplaceSystems | Note drawers are optional, but nice to have
| 3123HD | Drawer brackets 36" heavy duty | 1 | each | 15.9 | 15.9 | WorkplaceSystems
831.8
Miscellaneous |  | decide for yourself if you need it or not…
| WB190330 | Power Strip 6 Outlets And Surge Protector W/ 15' Cord | 0 | each | 49.95 | 0 | Global Industrial | check out their selection | http://www.globalindustrial.com/gcs/group/2236/productsPerPicGroups.web#gridAnchor
| WB190320 | Power Strip 6 Outlets And Surge Protector, 6 ft cord | 0 | each | 44.95 | 0 | Global Industrial
| 332750 | Apollo Safety Belt With Ratchet, Black | 1 | each | 35.29 | 35.29 | Office Depot | to attach the PC to the table, resting it on the platform of the leg
| 934319 | Tripp Lite 25' Black Cat6 Gigabit Snagless Patch Cable, RJ-45 (M) - RJ-45 (M) - UTP - ( CAT 6 ) - black | 0 | each | 11.96 | 0 | CDW Government | 25ft long cable for FT
| 1013476 | Tripp Lite Tripp Lite 50’ Gray Cat6 Gigabit Snagless Patch Cable, 50ft, Patch cable - RJ-45 (M) - RJ-45 (M) - 50 ft - UTP - ( CAT 6 ) - molded, stranded, snagless, booted - gray | 0 | each | 21.94 | 0 | CDW Government | 50ft one if you prefer
Noise reduction boxes for the two rough pumps: |  |  | NOTE tell them that you have an ETD, and to include the appropriate lid for the small rough pump.
| SSH11TF | Noise Reduction Box SSH11TF 350x300x750 (WxHxD) | 3 | each | 1270 | 3810 | Sonation GmbH | http://www.sonation.de/pages/en/home.php
| OP1141 | Oil pan for SSH11TF | 3 | each | 240 | 720 | Sonation GmbH
| FK-KF25 | Flange Extender Kit consisting of elongation 100mm, gasket and clamp | 3 | each | 48 | 144 | Sonation GmbH
HPLC
nanoAcquity | you need an extra Ethernet card to connect to the PC that comes with the mass spectrometer, the PC comes with two cards, one dedicated for the mass spectrometer, and one for the internet, you'll need a third one for the nanoAcquity
| 568872 | NETGEAR GA311 Gigabit PCI Adapter (GA311NA) | 1 | each | 25.93 | 25.93 | CDW Government | Ethernet card for nanoAcquity |  | GA311NA
configure the card:
| IP address | 193.168.0.1
| Subnet mask | 255.0.0.0
| Default gateway | . . .
use the following DNS server
| Preferred DNS server | 128.95.120.1
| Alternate DNS server | 128.95.112.1
Fitting to connect the nanoAcquity sparging line to your house line
| SS-QC4-S-200 | SS Instrumentation Quick-Connect Stem w/o Valve, 0.08 Cv, 1/8 in. Swagelok Tube Fitting | 1 | each | 23.99 | 23.99 | Swagelok | for nanoAcquity sparging, this is what we use at UWPR, but you need to find the correct fitting for your house line
| Driver | get the Driver for the nanoAcquity from Thermo
| contact closure cable | from Thermo, usually part of your quote
Agilent | Drivers include in Xcalibur
Eksigent | download drivers from | http://www.eksigent/setup